A close up of a sign

Description automatically generated

# QUALITY CONTROL CHECKLIST FOR BIOLOGICAL INDICATOR TEST

Maintain this checklist in your BioRAFT Equipment tab for six years. Questions: Contact EH&S Biosafety Officer at (254) 710-2900.

|  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **Autoclave make/model:** | |  | | | | **Location**  **(building/room number):** | | |  | |
| **Lab/Facility name:** | |  | | | | **Principal Investigator/ Supervisor name:** | | |  | |
| **Person responsible for autoclave:** | |  | | | | **Phone number:** | | |  | |
| **Email:** | |  | | | | **Location of autoclave records:** | | |  | |
| **Year:** | | | **Biological Indicator** | | | | **Brand** | **Cycle Time** | | **Sterilization Temperature** |
| **Month** | **Day** | | **Autoclaved**  **Indicator** | | **Positive**  **Control** | |
| **January** |  | |  | |  | |  |  | |  |
| **February** |  | |  | |  | |  |  | |  |
| **March** |  | |  | |  | |  |  | |  |
| **April** |  | |  | |  | |  |  | |  |
| **May** |  | |  | |  | |  |  | |  |
| **June** |  | |  | |  | |  |  | |  |
| **July** |  | |  | |  | |  |  | |  |
| **August** |  | |  | |  | |  |  | |  |
| **September** |  | |  | |  | |  |  | |  |
| **October** |  | |  | |  | |  |  | |  |
| **November** |  | |  | |  | |  |  | |  |
| **December** |  | |  | |  | |  |  | |  |
|  | | | | | | | | | | |
| **INTERPRETATION OF RESULTS** | | | | | | | | | | |
| **Autoclaved Indicator** | | **Positive Control** | | **Meaning** | | | | | | |
| No growth | | Growth | | Sterilization of spores achieved. | | | | | | |
| Growth | | Growth | | Autoclave cycle failure; troubleshoot autoclave and get repairs if needed. Test  again with new biological indicators. | | | | | | |
| Growth | | No growth | | Operator error; test again with new biological indicators. | | | | | | |
| No growth | | No growth | | Indicator may be expired; test again with unexpired biological indicators. | | | | | | |

A close up of a sign

Description automatically generated

# BIOLOGICAL INDICATOR TEST INSTRUCTIONS

Monthly biological indicator monitoring is required for all autoclaves used for sterilization of biohazardous waste at UW. This test validates that the autoclave is sufficiently inactivating biological or infectious materials.

|  |  |
| --- | --- |
| **EQUIPMENT** | |
| Biological indicator | Use *Bacillus stearothermophilus* spore strips or ampoules with an average population of 104 to 106 organisms. |
| Supporting laboratory equipment as needed | May include (depending on spore product in use):   * Incubator or heat block * Refrigerator * Culture media * Glassware |
| Quality Control Checklist | Use to records results. |
| **PROCEDURE** | |
| 1. Place indicator in center of load. | |
| 2. Select cycle used for sterilization of biohazardous waste. The cycle must include a **minimum temperature of 121°C or 250°F for 30 minutes or longer**, depending on size and compaction of the load. The full cycle time will take 60-90 minutes | |
| 3. Record chamber temperature on the Quality Control Checklist. | |
| 4. Remove indicator when cycle is complete and load is cool enough to handle. | |
| 5. Follow manufacturer’s instructions for activating and incubating the indicator. Remember to incubate a non-autoclaved indicator as a positive control. | |
| 6. Record results on Quality Control Checklist. | |
| 7. If growth occurs in the processed indicator, the autoclave has not inactivated the spores. Immediately contact the person responsible for the autoclave. The autoclave must remain out of service until a successful biological indicator test is achieved. The responsible person must inform users not to use the autoclave and post a “Do Not Use” sign while the autoclave is out of service. | |
| 8. Maintain the Quality Control Checklist in the lab/facility for six years. | |